# MentoScope — Classifier Misclassification Diagnostic Report

Document ID: MS-DBG-001
Date: 12 July 2026
Subject: AOM sample misclassified as Cerumen Impaction — root cause analysis

# Executive Summary

A known AOM (Acute Otitis Media) sample image was misclassified as Cerumen Impaction, with the Cerumen class probability scoring approximately 9 percentage points higher than AOM. This report documents the systematic diagnosis of the root cause, the evidence gathered, and the fix applied.

# 1. Resize Consistency Check

The feature extraction function (extract_features in features.py) centralizes image resizing to IMG_SIZE = (128, 128) using cv2.INTER_AREA interpolation. Both the training script (train_classifier.py) and the inference path (mentoscope_app.py) import this function from the same file, so the resize dimensions and method are identical by construction.

Conclusion: Resize is NOT the cause. Both paths use (128, 128) via the same code.

# 2. Crop Geometry Mismatch — The Root Cause

The classifier model (classifier_ear.joblib) was trained using crops produced by localize_region(), which — with USE_REAL_MODEL=False — applies a fixed centre-60% stub box to every image regardless of content. However, the inference pipeline was switched to LOCALIZER_MODE='yolo', which produces bounding boxes from a trained YOLOv8 detector. These two crop geometries are different:

• Training crop: centre-60% stub box [e.g. (100, 100, 400, 400) on a 500×500 image]
• Inference crop: YOLO-detected box [e.g. (81, 1, 497, 441) — wider, includes more border]
• The model learned features on one geometry but receives a different geometry at inference.

Conclusion: The crop mismatch between training (stub) and inference (YOLO) is the root cause.

# 3. Metadata Confirmation

The model's metadata file (models/meta_ear.json) confirms the mismatch:

• localizer: "stub" (training used the centre box)
• use_real_model: false
• n_samples: 3014
• trained_at: 2026-07-10T08:05:44

The inference path now uses LOCALIZER_MODE=yolo, which does not match the localizer field in the metadata.

# 4. Direct Feature-Extraction A/B Test

To confirm the diagnosis, the same AOM and Cerumen images were classified using both crop methods through the identical feature extraction + model pipeline. The results definitively show the crop geometry causing the prediction flip:

| Image | STUB crop (training geometry) | YOLO crop (inference geometry) | Verdict |
| --- | --- | --- | --- |
| AOM_000.jpg | AOM 0.83 ✓ | Cerumen 0.38 ✗ (AOM dropped to 0.24) | FLIPPED |
| AOM_001.jpg | AOM 0.63 ✓ | AOM 0.30 ✓ (confidence halved) | Degraded |
| AOM_002.jpg | AOM 0.89 ✓ | AOM 0.40 ✓ (confidence halved) | Degraded |
| Cerumen CI_000 | Cerumen 0.99 ✓ | Cerumen 0.90 ✓ | Correct |
| Cerumen CI_001 | Cerumen 1.00 ✓ | Cerumen 0.85 ✓ | Correct |
| Cerumen CI_002 | Cerumen 0.99 ✓ | Cerumen 0.68 ✓ | Correct |

Key finding: AOM_000 flips from a confident correct AOM (0.83) under the stub crop to a WRONG Cerumen (0.38 > AOM 0.24) under the YOLO crop. The model CAN separate AOM from Cerumen — it just receives different pixels than it was trained on.

# 5. Not a Training-Time Confusion Issue

Cerumen images classify correctly under both crop geometries (0.99 stub, 0.85-0.90 YOLO), proving the HOG + colour-histogram features DO separate AOM from Cerumen during training. The confusion only appears at inference when the crop geometry changes — this is a pipeline bug, not a feature or data limitation.

# 6. Fix Applied

train_classifier.py was updated to crop training images using the SAME LOCALIZER_MODE the app uses at inference:

• LOCALIZER_MODE == 'yolo' → crops with localize_by_yolo_ear() (YOLO box)
• LOCALIZER_MODE == 'texture' → crops with localize_by_texture()
• LOCALIZER_MODE == 'stub' → legacy localize_region() centre box

A new function _crop_box_for_training() in train_classifier.py resolves the active localizer and applies it consistently. The metadata now records the actual LOCALIZER_MODE instead of the stale 'stub' value.

Required action: Retrain the ear classifier (python train_classifier.py --mode ear) so the model learns on YOLO-cropped features that match inference.

# 7. Honest Scope Note

This fix resolves the crop-geometry mismatch between training and inference. However, the YOLO ear detector (yolo_ear.pt) was trained on synthetic images (real ear photos pasted onto solid-colour/gradient/noise backgrounds). Training mAP50 was 0.995 on a synthetic validation set — this proves the synthetic-label technique works, NOT that the model is 99.5% accurate on real endoscope frames. Real-frame localization and classification performance remain unvalidated until tested on real otoscopy footage.